## Slide 1
JoTouch
AI-Driven Prosthetic Hand
TEAM: Issa · Mohammad ·Abdallah
SUPERVISOR: Dr. Mohammad Khrisat
Al-Balqa Applied University
Notes: Open. JoTouch — an AI-driven anthropomorphic prosthetic hand, graduation project at Al-Balqa Applied University. Team of three, seven-month build. This is a Phase 1 catch-up: the digital twin is complete and running.

## Slide 2
OVERVIEW
02 / 22
The What
Anthropomorphic prosthetic hand
AI-driven via Force Myography
ROS 2 + custom firmware
↳ RViz · one → two → three → four → five
Notes: What JoTouch is in one breath: an anthropomorphic hand driven by AI from forearm muscle signals — Force Myography — running on ROS 2 with custom firmware. Seven-month graduation scope. The loop on the right is the live RViz simulation cycling gestures.

## Slide 3
01
SECTION 01 · PHASE 1 COMPLETE
Mechanical
Mechanical Design · URDF
Notes: Section one. This is the half of Phase 1 about the physical machine — the mechanical design, the CAD, and turning that CAD into a simulatable URDF model. All complete.

## Slide 4
PHASE 1 · MECHANICAL
04 / 22
1 Motor → 3 Joints per finger
Master–slave kinematics — MCP drives, PIP & DIP follow via four-bar linkage
F1404 BLDC
MCP
90°
45 mm proximal
PIP
100°
27 mm middle
DIP
70°
18 mm distal
FOUR-BAR LINKAGE · ANATOMICAL COUPLING
ONE motor per finger drives the MCP — PIP & DIP follow. No per-joint actuators.
5
motors, not 15
Notes: This is the key architectural decision. One motor drives the MCP of each finger; the PIP and DIP joints follow mechanically through a four-bar linkage — anatomical coupling at 90, 100, and 70 degrees at full curl. So three joints per finger move from a single actuator. Five motors instead of fifteen, while keeping a natural curl.

## Slide 5
PHASE 1 · MECHANICAL
05 / 22
CAD & Constraints
Total mass
290 g
Phalanges (Fibonacci)
45 / 27 / 18 mm
Construction
PA12 nylon
Gear train
brass + POM
Reduction / finger
400 : 1
Notes: The CAD locks the physical constraints. Total hand mass is 290 grams measured in CAD. Phalange lengths follow a Fibonacci progression — 45, 27, 18 millimeters. Printed in PA12 nylon, with a brass and Delrin gear train running a 400-to-1 reduction per finger. Seven actuated degrees of freedom.

## Slide 6
PHASE 1 · MECHANICAL
06 / 22
CAD to URDF
SOURCE
SolidWorks
→
EXPORT + SCRIPT
SW2URDF
→
SIMULATION
ROS 2 · RViz
16
links
14
joints
8
mimic relationships
Notes: Getting from CAD to a simulatable model. We export from SolidWorks with SW2URDF, then run custom Python post-processing — flipping joint axes, fixing mesh paths, and injecting the mimic joints that encode the coupling. The result: 16 links, 14 joints, and 8 mimic relationships, all driving the digital twin.

## Slide 7
02
SECTION 02 · HARDWARE
Electrical
Split compute · 8-channel FMG sensor array

## Slide 8
System Architecture
08 / 22

## Slide 9
ARCHITECTURE
09 / 22
Brain & Spine
Split compute · ROS 2 Jazzy on Ubuntu 24.04
BRAIN
Raspberry Pi Zero 2 W
ML inference
TensorFlow Lite
micro-ROS
◂ ──────── ▸
USB serial
SPINE
Teensy 4.1
Real-time motor control
600 MHz · ample headroom SimpleFOC · closed loop
Notes: The compute is split into a brain and a spine. The Raspberry Pi Zero 2 W is the brain — it runs ML inference. The Teensy 4.1 is the spine — hard real-time motor control at 600 MHz. They talk over micro-ROS on a USB serial link. This split keeps inference latency off the control loop.

## Slide 10
04 · CODING THE TEENSY & RASPBERRY PI
Two controllers, one real-time loop
IN PROGRESS
Teensy 4.1 real-time control (C++)
Reads the sensors
captures the muscle signal
Runs the on-board mapping
base weights from our data, updated per patient
Drives the motors
moves the joints smoothly + proportional
Reads the encoders
feeds back true motion for control & training
Raspberry Pi AI prediction (Python)
Runs the AI model
classifies gesture + regresses joint angles
Sends commands when confident
acts only above the confidence target
Handles online training
updates each patient's weights
Coordinates the two units
high-level brain to the Teensy's spine
How they talk to each other
Teensy
sensor + encoder stream →
Raspberry Pi
predicts intent
Teensy
← target motion commands
Goal: smooth, real-time response end-to-end · commands sent only when the model is confident.
coded
to wire up
JoTouch — Technical Brief · Sensing, Control & AI
10 / 22

## Slide 11
Technical Specifications
Performance Targets
Degrees of Freedom
Multi-Degree of Freedom
Control Latency
< 100ms target
Intent Accuracy
> 90% target
Operating Time
8+ hours daily use
Weight Target
< 700g
System Specifications
Processing
Teensy 4.1 + RPi Zero 2 W
Communication
UART / micro-ROS
Power
2S 18650 Li-ion (7.4V)
Sensing
FMG , BMP388 + Temp Input
Framework
ROS 2 + TensorFlow Lite
Sensory Feedback System
Pressure sensors on fingertips → Servo-driven rubber band on residual limb | Temperature sensors → Heating pad feedback
11 / 22

## Slide 12
01 · WHERE WE READ THE SIGNAL
8 sensors over the forearm muscles
PLANNED
forearm
cross-section
S1
S2
S3
S4
S5
S6
S7
S8
Flexors — close
Extensors — open
8-sensor ring · proximal forearm
Sensor → muscle → what it controls
S1
Brachioradialis (BR) — wrist side movement + partial flexion
S2
Ext. Carpi Radialis (ECR) — wrist extension + radial (toward thumb)
S3
Extensor Digitorum (ED) — opens fingers 2–5 — opens the hand
S4
Ext. Carpi Ulnaris (ECU) — wrist extension + ulnar (toward little)
S5
Flexor Carpi Ulnaris (FCU) — wrist flexion + ulnar deviation
S6
Flexor Digit. Superficialis (FDS) — flex middle / ring / little — grip
S7
FDS (index region) — flex index finger — pinch / precision
S8
Flexor Carpi Radialis + FPL — wrist flexion + thumb — pinch
Reading the right signal
Sensors sit over the muscle bellies on the proximal forearm — about 5–6 cm below the elbow — where the FMG signal is strongest. Flexor muscles on the palm side close the fingers and wrist; extensor muscles on the back open them. Eight evenly-spaced sensors give full-circle coverage, and the model learns the natural overlap between neighbouring muscles (crosstalk) on its own — no muscle has to be isolated perfectly.
JoTouch — Technical Brief · Sensing, Control & AI
12 / 22

## Slide 13
03
SECTION 03 · PHASE 1 COMPLETE
ROS 2 Pipeline & Software
7 packages · synthetic FMG → brain → kinematics → RViz
Notes: Section two — the software half of Phase 1. Seven ROS 2 packages built from scratch form the pipeline that drives the digital twin, end to end.

## Slide 14
SOFTWARE · COLLECTION UI
14 / 22
FMG Data Collection Workbench
Browser operator console — Setup → Tests → Collect → Summary
01 SETUP — session IDs, protocol, LED sync ROI, live FSR preview
02 TESTS — 8 mandatory pre-collection checks with live cueing

## Slide 15
SOFTWARE · COLLECTION UI
15 / 22
Collecting a Session
Cue-driven recording, real-time quality gating, BIDS-structured output
03 COLLECT — cue card, camera, FSR signals, run log
04 SUMMARY — run tally, quality report, full session log

## Slide 16
02 · AI MODEL
JoTouchModel — classification and regression
BENCHMARKED
Not just classification: the model predicts which gesture (classification) and the exact joint angles (regression) → proportional, natural control.
How the model reads the signal
Windowed FMG input
a short slice of sensor history
MobileNetV2 blocks
extract features from the signal
TCN (temporal conv.)
model how the signal evolves in time
Regularised shared layer
robust, compact representation
Two heads — Classification + Regression
gesture + continuous joint angles
Why these choices
Why classification + regression
Classification picks the intended gesture; regression outputs the exact joint angles. Together they give smooth, proportional motion instead of fixed on/off grips.
Why MobileNetV2
Lightweight and efficient — built for edge devices — so it runs on the Raspberry Pi within our power and latency budget, with far fewer parameters than a full CNN.
Why a TCN (not an RNN/LSTM)
Convolutions over time capture the muscle pattern as it unfolds, run fast in parallel, and avoid the unstable gradients of recurrent nets — better for real-time control.
Why on-device + confidence gating
Local inference removes network delay and protects privacy; the hand acts only when the model is confident, keeping motion safe and stable.
JoTouch — Technical Brief · Sensing, Control & AI
16 / 22

## Slide 17
03 · ONLINE PER-PATIENT TRAINING
The hand learns each patient over time
PLANNED
Sensor → motor weights
① Initial fixed weights trained on our dataset
② Per-patient updated online for each user over time
How we get the signals
Input — intent
The forearm FMG sensors capture what the patient is trying to do.
Output — ground truth
The motor encoders report the real joint positions, giving a matched (intent → motion) pair every moment — no extra lab equipment.
Supervised calibration
In a guided session the patient performs prescribed movements, so the collected pairs are trustworthy labels for learning.
The personalisation loop
1 · Use / calibrate
FMG in · encoder angles out
2 · Learn online
refine on this patient's data
3 · Accuracy check
does it reach our target?
4 · Update weights
push refined weights to Teensy
once target accuracy is met
Why online, per-patient
Anatomy and muscle activation differ per person and drift over time, so a fixed model degrades. Encoders supply ground truth continuously, letting the system keep improving itself.
JoTouch — Technical Brief · Sensing, Control & AI
17 / 22

## Slide 18
05 · TRAINING DATA & RESULTS
Trained and tested on open-source FMG datasets
OPEN DATA
To be precise: we have not run our own human trials yet. JoTouchModel was trained end-to-end on three published open-source FMG datasets — 37,500 windows — and scored on a 15% test set it never saw while training.
The three datasets
Zenodo MDPI FMG · 17 subjects
Zakia & Menon, MDPI Data 2022 · 5 stage motions 12,500 windows · 8 channels × 20 steps (200 ms)
Exoskelebox FMG benchmark · 20 subjects
arXiv:2007.14918 · 6 gestures × 3 wrist orientations 15,000 windows · 8 forearm FSR @ ~100 Hz
PLoS ONE 2025 FMG · 27 participants
Young et al., e0321319 · 4 grasps × 5 load levels 10,000 windows · 8 channels @ 100 Hz
Accuracy on the untouched test set
Dataset
Test accuracy
Latency
Epochs
Zenodo (5 motions)
86.78%
3.57 ms
13
Exoskelebox (6 gestures)
70.92%
3.07 ms
40
PLoS ONE (4 grasps)
42.61%
2.06 ms
12
One model, small enough for the hand itself
65.8 K parameters · 257 KB · 2–3.6 ms per window on a plain CPU. The window is 200 ms and refreshes every 50 ms, so prediction uses only a few percent of the time we have — the latency budget is not the bottleneck.
How these numbers were produced
The data is public and cited, so anyone can reproduce this. One model per dataset — the label sets mean different things, so merging them into a single classifier would be dishonest. Identical pipeline for all three: AdamW + OneCycleLR + gradient clipping + early stopping (patience 8), up to 40 epochs, batch 64, lr 0.001, split 70% train / 15% validation / 15% untouched test.
JoTouch — Technical Brief · Sensing, Control & AI
18 / 22

## Slide 19
06 · HOW WELL IT GENERALISES
The same subject vs. an unseen subject
HONEST LIMITS
The limit we do not hide: accuracy measured on a subject the model has never seen is far lower than accuracy on the subjects it trained on. We tested this on purpose, with leave-one-subject-out cross-validation.
Identical training, two ways of splitting the data
Same subjects (random split)
Unseen subject (leave-one-subject-out)
Zenodo
17 subjects · 5 classes · chance 20%
96.8%
35.4% ±9.6
Exoskelebox
20 subjects · 6 classes · chance 17%
64.5%
21.2% ±2.7
PLoS ONE
27 subjects · 4 classes · chance 25%
50.6%
29.4% ±1.5
Both bars come from the identical 12-epoch protocol on identical data — only the split differs.
Why — and what we do about it
Why the drop happens
FMG reads the muscle bulging under the skin. Forearm size, fat layer, band tightness and the exact rotation of the band differ from person to person, so the same gesture produces a different 8-channel pattern on a new arm.
Why this is not a dead end
This is exactly the case for per-patient training (slide 3). The motor encoders supply ground truth for free, so the hand calibrates on its own user instead of depending on a stranger's forearm.
Regression too: R² 0.85 → 0.03 on an unseen subject (Zenodo).
What we claim — and what we do not
We claim a real-time model that learns an individual subject's FMG signal well on published open data, and runs inside the hand's time and power budget. We do not claim a plug-and-play model that works on an unseen arm out of the box — no published FMG model does — and we have not yet collected data from our own patients. The per-patient calibration step is where that happens, and it is the engineering answer to this limit, not a workaround for it.
JoTouch — Technical Brief · Sensing, Control & AI
19 / 22

## Slide 20
PHASE 1 · SOFTWARE
20 / 22
ROS 2 Dataflow
INPUT
COMPUTE
OUTPUT
/jotouch_fmg
synthetic FMG @ 200 Hz
/jotouch/fmg
───▸
/jotouch_brain
ML inference · dummy model
/jotouch/joint_command
───▸
/jotouch_kinematics
7 DOF → 14 joint angles
/joint_states
▼
/jotouch_spine_sim
fake Teensy · 1st-order lag
/robot_state_pub
forward kinematics → /tf
/tf
───▸
RViz
3D visualization
Notes: The live dataflow. Synthetic FMG at 200 hertz feeds the brain node — currently a dummy linear model. The brain emits 7 joint curls; kinematics expands those to 14 joint angles using the master-slave coupling. From there it splits: the spine simulator stands in for the Teensy, and robot-state-publisher drives RViz. Input, compute, output — color-coded.

## Slide 21
PHASE 1 · SOFTWARE
21 / 22
Tools & Stack
ROS 2 Jazzy
MIDDLEWARE
Python 3.12
ROS 2 NODES
C++
TEENSY FIRMWARE
PlatformIO
BUILD / FLASH
SimpleFOC
FIELD-ORIENTED CTRL
micro-ROS
BRAIN ↔ SPINE
TF Lite Micro
DEPLOY TARGET
Obsidian
DOCUMENTATION
Notes: The toolchain. ROS 2 Jazzy on the software side, Python for the nodes, C++ on the Teensy through PlatformIO, SimpleFOC for field-oriented motor control, micro-ROS bridging the spine, TensorFlow Lite Micro as the deployment target for the model, and Obsidian holding all the project documentation.

## Slide 22
JOTOUCH //
PHASE 1 · COMPLETE
Questions?
Kinematics · ROS 2 stack · the road to hardware
TEAMIssa · Abdullah · Mohammad
SUPERVISORDr. Mohammad Khrisat
Al-Balqa Applied University · 2026
Notes: That's Phase 1 — complete and validated. Happy to go deeper on the kinematics, the ROS 2 stack, or the road to hardware. Thank you. Questions?