## Slide 1
01 · WHERE WE READ THE SIGNAL
8 sensors over the forearm muscles
PLANNED
forearm
cross-section
S1
S2
S3
S4
S5
S6
S7
S8
Flexors — close
Extensors — open
8-sensor ring · proximal forearm
Sensor → muscle → what it controls
S1
Brachioradialis (BR) — wrist side movement + partial flexion
S2
Ext. Carpi Radialis (ECR) — wrist extension + radial (toward thumb)
S3
Extensor Digitorum (ED) — opens fingers 2–5 — opens the hand
S4
Ext. Carpi Ulnaris (ECU) — wrist extension + ulnar (toward little)
S5
Flexor Carpi Ulnaris (FCU) — wrist flexion + ulnar deviation
S6
Flexor Digit. Superficialis (FDS) — flex middle / ring / little — grip
S7
FDS (index region) — flex index finger — pinch / precision
S8
Flexor Carpi Radialis + FPL — wrist flexion + thumb — pinch
Reading the right signal
Sensors sit over the muscle bellies on the proximal forearm — about 5–6 cm below the elbow — where the FMG signal is strongest. Flexor muscles on the palm side close the fingers and wrist; extensor muscles on the back open them. Eight evenly-spaced sensors give full-circle coverage, and the model learns the natural overlap between neighbouring muscles (crosstalk) on its own — no muscle has to be isolated perfectly.
JoTouch — Technical Brief · Sensing, Control & AI
1 / 6

## Slide 2
02 · AI MODEL
JoTouchModel — classification and regression
BENCHMARKED
Not just classification: the model predicts which gesture (classification) and the exact joint angles (regression) → proportional, natural control.
How the model reads the signal
Windowed FMG input
a short slice of sensor history
MobileNetV2 blocks
extract features from the signal
TCN (temporal conv.)
model how the signal evolves in time
Regularised shared layer
robust, compact representation
Two heads — Classification + Regression
gesture + continuous joint angles
Why these choices
Why classification + regression
Classification picks the intended gesture; regression outputs the exact joint angles. Together they give smooth, proportional motion instead of fixed on/off grips.
Why MobileNetV2
Lightweight and efficient — built for edge devices — so it runs on the Raspberry Pi within our power and latency budget, with far fewer parameters than a full CNN.
Why a TCN (not an RNN/LSTM)
Convolutions over time capture the muscle pattern as it unfolds, run fast in parallel, and avoid the unstable gradients of recurrent nets — better for real-time control.
Why on-device + confidence gating
Local inference removes network delay and protects privacy; the hand acts only when the model is confident, keeping motion safe and stable.
JoTouch — Technical Brief · Sensing, Control & AI
2 / 6

## Slide 3
03 · ONLINE PER-PATIENT TRAINING
The hand learns each patient over time
PLANNED
Sensor → motor weights
① Initial fixed weights trained on our dataset
② Per-patient updated online for each user over time
How we get the signals
Input — intent
The forearm FMG sensors capture what the patient is trying to do.
Output — ground truth
The motor encoders report the real joint positions, giving a matched (intent → motion) pair every moment — no extra lab equipment.
Supervised calibration
In a guided session the patient performs prescribed movements, so the collected pairs are trustworthy labels for learning.
The personalisation loop
1 · Use / calibrate
FMG in · encoder angles out
2 · Learn online
refine on this patient's data
3 · Accuracy check
does it reach our target?
4 · Update weights
push refined weights to Teensy
once target accuracy is met
Why online, per-patient
Anatomy and muscle activation differ per person and drift over time, so a fixed model degrades. Encoders supply ground truth continuously, letting the system keep improving itself.
JoTouch — Technical Brief · Sensing, Control & AI
3 / 6

## Slide 4
04 · CODING THE TEENSY & RASPBERRY PI
Two controllers, one real-time loop
IN PROGRESS
Teensy 4.1 real-time control (C++)
Reads the sensors
captures the muscle signal
Runs the on-board mapping
base weights from our data, updated per patient
Drives the motors
moves the joints smoothly + proportional
Reads the encoders
feeds back true motion for control & training
Raspberry Pi AI prediction (Python)
Runs the AI model
classifies gesture + regresses joint angles
Sends commands when confident
acts only above the confidence target
Handles online training
updates each patient's weights
Coordinates the two units
high-level brain to the Teensy's spine
How they talk to each other
Teensy
sensor + encoder stream →
Raspberry Pi
predicts intent
Teensy
← target motion commands
Goal: smooth, real-time response end-to-end · commands sent only when the model is confident.
coded
to wire up
JoTouch — Technical Brief · Sensing, Control & AI
4 / 6

## Slide 5
05 · TRAINING DATA & RESULTS
Trained and tested on open-source FMG datasets
OPEN DATA
To be precise: we have not run our own human trials yet. JoTouchModel was trained end-to-end on three published open-source FMG datasets — 37,500 windows — and scored on a 15% test set it never saw while training.
The three datasets
Zenodo MDPI FMG · 17 subjects
Zakia & Menon, MDPI Data 2022 · 5 stage motions 12,500 windows · 8 channels × 20 steps (200 ms)
Exoskelebox FMG benchmark · 20 subjects
arXiv:2007.14918 · 6 gestures × 3 wrist orientations 15,000 windows · 8 forearm FSR @ ~100 Hz
PLoS ONE 2025 FMG · 27 participants
Young et al., e0321319 · 4 grasps × 5 load levels 10,000 windows · 8 channels @ 100 Hz
Accuracy on the untouched test set
Dataset
Test accuracy
Latency
Epochs
Zenodo (5 motions)
86.78%
3.57 ms
13
Exoskelebox (6 gestures)
70.92%
3.07 ms
40
PLoS ONE (4 grasps)
42.61%
2.06 ms
12
One model, small enough for the hand itself
65.8 K parameters · 257 KB · 2–3.6 ms per window on a plain CPU. The window is 200 ms and refreshes every 50 ms, so prediction uses only a few percent of the time we have — the latency budget is not the bottleneck.
How these numbers were produced
The data is public and cited, so anyone can reproduce this. One model per dataset — the label sets mean different things, so merging them into a single classifier would be dishonest. Identical pipeline for all three: AdamW + OneCycleLR + gradient clipping + early stopping (patience 8), up to 40 epochs, batch 64, lr 0.001, split 70% train / 15% validation / 15% untouched test.
JoTouch — Technical Brief · Sensing, Control & AI
5 / 6

## Slide 6
06 · HOW WELL IT GENERALISES
The same subject vs. an unseen subject
HONEST LIMITS
The limit we do not hide: accuracy measured on a subject the model has never seen is far lower than accuracy on the subjects it trained on. We tested this on purpose, with leave-one-subject-out cross-validation.
Identical training, two ways of splitting the data
Same subjects (random split)
Unseen subject (leave-one-subject-out)
Zenodo
17 subjects · 5 classes · chance 20%
96.8%
35.4% ±9.6
Exoskelebox
20 subjects · 6 classes · chance 17%
64.5%
21.2% ±2.7
PLoS ONE
27 subjects · 4 classes · chance 25%
50.6%
29.4% ±1.5
Both bars come from the identical 12-epoch protocol on identical data — only the split differs.
Why — and what we do about it
Why the drop happens
FMG reads the muscle bulging under the skin. Forearm size, fat layer, band tightness and the exact rotation of the band differ from person to person, so the same gesture produces a different 8-channel pattern on a new arm.
Why this is not a dead end
This is exactly the case for per-patient training (slide 3). The motor encoders supply ground truth for free, so the hand calibrates on its own user instead of depending on a stranger's forearm.
Regression too: R² 0.85 → 0.03 on an unseen subject (Zenodo).
What we claim — and what we do not
We claim a real-time model that learns an individual subject's FMG signal well on published open data, and runs inside the hand's time and power budget. We do not claim a plug-and-play model that works on an unseen arm out of the box — no published FMG model does — and we have not yet collected data from our own patients. The per-patient calibration step is where that happens, and it is the engineering answer to this limit, not a workaround for it.
JoTouch — Technical Brief · Sensing, Control & AI
6 / 6